## Slide 1
Partially Polarized Calibration Stimulus
Gerard van Hartena, Frans Snika, Christoph Kellera,
Jeroen Rietjensb, Martijn Smitb
aLeiden University, The Netherlands
bSRON Netherlands Institute for Space Research
Calibration Workshop of the MP1104 COST Action
January 2013
Zurich, Switzerland

## Slide 2
Atmospheric aerosols
Size
Shape
Chemical composition

## Slide 3
Measurements
Stokes I, Q, U
Multiple wavelengths
Multiple scattering angles

## Slide 4
Aerosol cooling & climate sensitivity mixed.
Disentangle in 20 years:
400-800 nm
30 viewing angles
σP=0.001+0.005×P
Instrument requirements
Knutti & Hegerl (2008)
Mishchenko et al. (2004)
Hasekamp & Landgraf (2007)
Notes: Climate sensitivity: warming effect of CO2
Polarimetric accuracy for refractive index requirement

## Slide 5
190 mm
100 mm
60 mm
SPEX prototype
spectral range: 400-750 nm
spectral resolution (flux): 1 nm
spectral resolution (polarization): 20 nm
volume: 1.1 dm3
mass: 1.0 kg

## Slide 6
SPEX prototype
100 mm
60 mm
spectral range: 400-750 nm
spectral resolution (flux): 1 nm
spectral resolution (polarization): 20 nm
volume: 1.1 dm3
mass: 1.0 kg
190 mm

## Slide 7
SPEX polarimetry
DoLP variation
AoLP variation
Snik et al. (2009)

## Slide 8
Polarimeter calibration
Rotating polarizer
P=1
σP=0.0001 (accuracy)
2. Tilted glass plate
0.05<P<0.40
σP=0.01 (stability)
STOTA calibration stimulus
0<P<0.40
σP=0.0005 (stability + accuracy)
Notes: 1.
Expensive polarizer -> high extinction ratio: by design
Polarimeters often measure very low P -> how does it behave? What is the zero-point, is the response linear?
Accuracy: we just don’t know / cannot promise anything below sP=0.0001
2.
Glass plate partial polarizer -> lamp polarization leaks through; polarization wavelength dependent
sensitive to lamp polarization fluctuations -> stability
3.
State-of-the-art stimulus
P down to 0
more

## Slide 9
Current SPEX calibration
Just a tilted glass plate
2×10-4 DoLP sensitivity
Notes: MM model – free parameters: n(glass), lamp polarization

## Slide 10
Stimulus philosophy
Depolarize to P<0.0005
Polarize σP<0.0005
Verify σP~0.0005
Mahler & Chipman (2011)
Notes: Depolarization:
Get low P
Be insensitive to lamp polarization fluctuations
Verification more accurate
Polarization:
0<P<40%
Tilt accuracy -> sP<0.0005
Verification:
- this in the end defines the stimulus’ accuracy

## Slide 11
Preliminary design
Lamp P~0.05
Fiber ×0.1???
Integrating sphere × <0.005
Liquid crystal scrambler (MTR) × <0.01
Crossed glassplates in gimbal mount
Slow lenses (P~0.0005 @ 2° refraction -> f~300 mm)
Notes: Fiber TBD
Very bright light source
Fiber coupling with integrating sphere
MTR: multi-twist retarder -> achromatic half-wave retarder; retarder axis random orientation every 10 micron
Different glass plate materials + coatings for different DoLP ranges / DoLP resolution / lambda ranges
AR coating on 1 side to prevent multiple reflections
2 crossed glass plates to cancel beam shift

## Slide 12
Multi-twist retarder
Self-aligning liquid crystals
Achromatic @ 3 layers
Spatial resolution ~ 10 μm
Komanduri et al. (2013)
Notes: - Substrate with pattern on 1 side
-

## Slide 13
Mahler & Chipman (2011)
Challenges
Pupil alignment
  Oversize
  Make uniform
  Measure map
  σP=0.00005 / mm
Lens polarization
Lens stress birefringence
Straylight / reflections
Verification @ σP=0.0005

## Slide 14
Chicken / egg
Model
  All separate components
  Total system – sensitivity analysis
Measure
  Imaging polarimetry pupil
  Different pupil alignments
Confidence: understanding + P~0.0001 measurement
Stimulus verification

## Slide 15
SPEX spectral polarization modulation
Snik et al. (2009)